# MRCO Commissioning Shift Plans (X-10168):

### Checkout Goals (from commissioning sheet TMO SCRF Commissioning V2):

### Initial Time Accounting

| Activity Name | Activity Objective | Estimated Beamtime (Hrs) | Beam Parameters | Note |
| --- | --- | --- | --- | --- |
| Define focus with WFS |  | 2 | 400 eV |  |
| Align Spectrometer | Verify X,Y position of XFEL focus: MRCO XY-scan to find the center of beam with respect to the centering ring. 0 retardation. (N2) | 2 | 950 eV | Sample is Ne. Look at all the ToFs to make sure we can center the beam. |
| Retardation Optimization | Ensure homegeneity of retardation voltages | 6 | 950 eV | Sample is Ne. Set up the lens cans of the spectrometers for different retardations (5 retardations). Check that the retardation is consistent across the ToFs. |
| Photon Energy/Retardation Scan | Characterize resolution by scanning retardation and electron kinetic energy | 3 | 415 - 1410 eV | Retardation scan of NNO for different photon energies & different retardations: hf = Ip + (5 [Vr=0], 10 [Vr=0,5], 20 [Vr=0,15], 50 [Vr=0,45], 100 [Vr=0,50,95], 200 [Vr=0,150,100,195], 500 [Vr=0,400,450,495], 1000 [Vr=0,500,900,950,995]) FZP = N, O, 613, Ne, 1200. This also gives measurement of beta parameters |
| Spectral Reconstruction | Correlate MRCO spectra with FZP spectrum | 4 | 950 eV (& if time, 570) | FZP/MRCO correlation measurements for spectral reconstruction Redo the above with a different hit rate (as a factor of 1, 2, 4, 8, 16* per feature). FZP/MRCO correlation measurements for spectral reconstruction in current mode Start with Ne & 950 eV, if time do CO2 as well |
| Test Current Mode | Ensuring that the spectrometer is tuned for high signal rate mode, and validate MCP performance in cases of increased depletion | 12 | 950 eV | Sample is Ne. We need to make sure that the MCP detectors are working at the non-linear gain 5, 10, 25, 50, 75, 100, 150, 200 mV signal level (without amplifier). Repeat 5, 10, 15, 25 after adding amplifiers. |
| Spectral Reconstruction in Current Mode |  | 4 |  | See above |
| Verify position of XFEL focus: | MRCO Z-scan to find nonlinear signal (DCH in Ne) | 6 | 960 eV (if this is a struggle, we use Ar @ 350 eV) | If the non-linear is difficult to find, then consider using the blurring effect of space charge to identify the tightest focus. |
| Laser spatial overlap (paddle) |  | 6 |  | IR/x-ray overlap |
| Laser temporal overlap (paddle) |  | 6 |  | IR/x-ray overlap |
| Laser spatiotemporal overlap (gas) |  | 6 | 400 eV | IR-pumped NNO RAM |
| More Beta Scans |  | 6 |  | N2, CO2, Ne |
| Total Hours |  | 61 |  |  |

### Detailed Shift Plan

#### Common tasks in all the shift:

- Check for KB1 transmission (TMO)

- Sample paddle repeatability (MRCO)

- Bias voltage monitoring - online plots. (MRCO)

- FEX monitoring - online plots (?). (MRCO)

- Data checking

- MRCO XY scan. (MRCO)

Shift Coverage_X10168

#### Shift 1:

- Bring beam down TMO beamline and check timing of imagers 
(timing should be fine, but just be mindful timing could have shifted)

  - Run beamline transmission script

- Ramp up of MRCO detectors (no retardation, 400 eV)

  - Check relative timing (XPM-2 vs XPM-4 should be ~1.3 us)

  - Align MRCO in X/Y plane (no retardation)

    - Measure AM electrons for different XY positions, and optimize on uniformity.

      - 400 eV - Ar -> L-MM (200 eV)

      - Optimize the relative yield between opposite ToFs.

    - Also check the flight time for the different ToFs (we can see if this shows something)

- We need to be above 1000 Hz to continue.

- MCP gain scan (with amplifier): bias scan: 1100:1500:50 V (in counting mode) 
(1 min each @ 1000 hz)

- Ramp MRCO Retardation (400 eV)
Adjust the retardation while looking at the angular distribution of Ar L-MM (200 eV).

  - 50 V, 100V, 150 V, 175 V, 190 V, 195 V.

- Setup energy scan PV with ACR.

- Spectrometer checkout Ming-Fu Lin (400 eV)

  - Check the timing of Piranha and Intensifier.

  - Align FZP w/ no attenuator (500 Hz)

  - Align Attenuator, then go to high rate.

- Swap to NNO

- Retardation scan (5-10 min at 1 kHz => 3 hrs)
hf = Ip + (5 [Vr=0], 10 [Vr=0,5], 20 [Vr=0,15], 50 [Vr=0,45], 100 [Vr=0,50,95], 200 [Vr=0,150,100,195], 500 [Vr=0,400,450,495], 1000 [Vr=0,500,900,950,995])

| FZP | Nitrogen (410 FZP) |  |  | Ti (460) | Oxygen (530) | Si (613) | 950 FZP | 1212 (fzp) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ret/hf | +5 | +10 | +20 | +50 | +100 | +200 | +500 | +800 |
| 0 | 417.5 | 422.5 | 432.5 | 462.5 | 512.5 | 612.5 | 912.5 | 1212.5 |
| 5 |  | 422.5 |  |  |  |  |  |  |
| 15 |  |  | 432.5 |  |  |  |  |  |
| 45 |  |  |  | 462.5 |  |  |  |  |
| 50 |  |  |  |  | 512.5 |  |  |  |
| 95 |  |  |  |  |  |  |  |  |
| 100 |  |  |  |  | 512.5 | 612.5 |  |  |
| 150 |  |  |  |  |  | 612.5 |  |  |
| 195 |  |  |  |  |  | 612.5 |  |  |
| 400 |  |  | 432.5 (V) |  | 512.5 (V) |  | 912.5 |  |
| 450 |  |  |  |  |  |  | 912.5 |  |
| 495 |  |  |  |  |  |  | 912.5 |  |
| 500 |  |  |  |  |  |  |  |  |

- Measure covariance of electron spectrum and photon spectrum

- Request 800 eV, 950 eV and see what the pulse energy (and count rate)

#### Shift 2:

- We need to be above 1000 Hz to continue.

- High-Rate scan (400 eV, Argon)

  - Adjust sample gas pressure (and x-ray flux) to have consistent signal in polarization ToFs.

  - MCP gain scan (no amplifier): bias scan: 1200:1850:50 V 
(1 min each @ 1000 hz)

- Setup Energy scan PV

- Spectral Reconstruction scans

  - Argon L-edge (400 eV)

  - Retardations: 0, 100, 125 V

  - Different hit rate (as a factor of 1, 2, 4, 8, 16* per feature).

  - Scanning photon energy across the FZP range.

  - 5 min at each retardation

#### Shift 3:

Pre-shift steps:

- Setup Argon gas line.

- Setup N2O in the gas cabinet (ready to be switched).

- Check all MRCO cameras.

- Make sure that the SMA pin in the sample paddle is connected to the scope.

- Check that the voltage ramper scripts have correct bias voltage for individual MCPs.

With Beam:

- Bring the beam down to the TMO beamline. Note down and point the beam in the same locations of IM2K4 and IM5K4. Day-to-day variations should be noted. (30 mins)

  - Check for KB1 transmission

  - Check for Sample paddle repeatability.

  - See if signal on the SMA tip. (Could try coarse timing at IP)

- Setup energy scan PV. (10 mins)

- Setup FZP at 400 eV. (30 mins)

  - If higher pulse energy > 40 uJ, use Target 2 of the attenuator.

- Ramp up of MRCO detectors and TOFs (no retardation, 400 eV) (10 mins)

- Measure background. (10 mins)

  - Start with 2-5% transmission (~10 uJ avg in the XGMD).

- Open Gas Needle valve connected to Argon line. (10 min)

- Do SQ1 XY scan around the current position. Delta X = +/- 1, Delta Y = +/- 1. Raster scanning by hand. (30 mins)

- Needle position scan in XYZ. (30 mins)

  - 100 V retardation.

  - Voltage scan at the needle. [-100 V to 100 V] ...

- Signal checking with and without amplifier (1 hour)

  - At most 4 TOF without amplifier.

  - Make sure FZP is in a safe state.

  - Transmission : 2% - 100%.

- Laser timing can be done here. (1 hour)

  - Spatial overlap on ATM

  - Temporal overlap at IP (could do this, optional)

  - Temporal overlap on ATM

- Adjust the retardation while looking at the angular distribution of Ar L-MM (200 eV). (30 mins @ 8.3 kHz, 2 mins /run) (~2-5% transmission)

  - 50 V, 100V, 150 V, 175 V, 190 V, 195 V, 200 V, 350 V at hv = [390 to 415 eV at 2 eV steps]

- Do SQ1 Z scan around the optimized XY positions. (30 mins)

- Move the focus downstream and repeat the retardation scan. Compare different Z (1 hour) Why? Do we see any non-linearity?

  - Z between +10 mm and -1 mm (make sure to account for Y due to beamline elevation. 0.04 deg)

- Spectral Reconstruction scans (2 hour)

  - Argon L-edge (400 eV).

  - Retardations: 0, 100, 125 V.

  - Different hit rate (as a factor of 1, 2, 4, 8, 16* per feature).

  - Scanning photon energy across the FZP range.

  - 5 min at each retardation

- Move on to NNO. (30 mins)

- Retardation scan (2 min at 8 kHz,  4 hours including setup of two FZPs)
hf = Ip + (5 [Vr=0], 10 [Vr=0,5], 20 [Vr=0,15], 50 [Vr=0,45], 100 [Vr=0,50,95], 200 [Vr=0,150,100,195], 500 [Vr=0,400,450,495],

  - Do photon energy/retardation scans around 400 eV FZP.

  - Change photon energy centered at 530 eV. Set up the FZP.

  - Do photon energy/retardation scans around 530 eV.

  - Change photon energy centered at 613 eV. Needs commissioning by Ming-Fu.

  - Do photon energy/retardation scans around 613 eV.

| FZP | Nitrogen (410 FZP) |  |  | Oxygen (530) | Si (613) |
| --- | --- | --- | --- | --- | --- |
| Ret/hf | +5 | +10 | +20 | +100 | +200 |
| 0 | 417.5 | 422.5 | 432.5 | 512.5 | 612.5 |
| 5 |  | 422.5 |  |  |  |
| 15 |  |  | 432.5 |  |  |
| 45 |  |  |  |  |  |
| 50 |  |  |  | 512.5 |  |
| 95 |  |  |  |  |  |
| 100 |  |  |  | 512.5 | 612.5 |
| 150 |  |  |  |  | 612.5 |
| 195 |  |  |  |  | 612.5 |
| 400 |  |  | 432.5 (V) | 512.5 (V) |  |
| 450 |  |  |  |  |  |

- Measure covariance of electron spectrum and photon spectrum

- Request 800 eV, 950 eV and see what the pulse energy (and count rate).

In parallel, if time allows:

- laser/x-ray spatial overlap on sample paddle

- Laser timing signal on sample paddle (SMA pin)

- ATM Setup

Contingencies and notes:

- Note down the bias voltages of the MCPs. The angular distribution of the Auger line should have a beta parameter = 0.05. If not, then evaluate the cause.

#### Shift 4

With Beam:

- Bring the beam down to the TMO beamline. Note down and point the beam in the same locations of IM2K4 and IM5K4. Day-to-day variations should be noted. (30 mins)

  - Check for KB1 transmission

  - Check for Sample paddle repeatability.

  - See if signal on the SMA tip. (Could try coarse timing at IP)

- Setup energy scan PV. (10 mins)

- Setup FZP at 400 eV. (30 mins)

  - If higher pulse energy > 40 uJ, use Target 2 of the attenuator.

- Signal checking with and without amplifier (1 hour)

  - At most 4 TOF without amplifier.

  - Make sure FZP is in a safe state.

  - Transmission : 2% - 40%.

- Resolution adjustment scan:

  - Opposing TOFs resolution. Tweak retardation to match the resolution and collection of each opposing TOF pairs.

    - Retardation of 175 , +/- 15 V retardation at a step of 5 V

- Laser timing can be done here. (1 hour)

  - Spatial overlap on ATM

  - Temporal overlap at IP (could do this, optional)

  - Temporal overlap on ATM

- Move the focus downstream and repeat the retardation scan. Compare different Z (1 hour) Why? Do we see any non-linearity?

  - Z between +10 mm and -1 mm (make sure to account for Y due to beamline elevation. 0.04 deg)

- Needle Scan? We need to look at the data for Retardation = 0.

  - We might want to repeat it with higher retardation, 190 V.

- Move on to NNO. (30 mins)

- Retardation scan (2 min at 8 kHz,  4 hours including setup of two FZPs)
hf = Ip + (5 [Vr=0], 10 [Vr=0,5], 20 [Vr=0,15], 50 [Vr=0,45], 100 [Vr=0,50,95], 200 [Vr=0,150,100,195], 500 [Vr=0,400,450,495],

  - Do photon energy/retardation scans around 400 eV FZP.

  - Change photon energy centered at 530 eV. Set up the FZP.

  - Do photon energy/retardation scans around 530 eV.

  - Change photon energy centered at 613 eV. Needs commissioning by Ming-Fu.

  - Do photon energy/retardation scans around 613 eV.

| FZP | Nitrogen (410 FZP) |  |  | Oxygen (530) | Si (613) |
| --- | --- | --- | --- | --- | --- |
| Ret/hf | +5 | +10 | +20 | +100 | +200 |
| 0 | 417.5 | 422.5 | 432.5 | 512.5 | 612.5 |
| 5 |  | 422.5 |  |  |  |
| 15 |  |  | 432.5 |  |  |
| 45 |  |  |  |  |  |
| 50 |  |  |  | 512.5 |  |
| 95 |  |  |  |  |  |
| 100 |  |  |  | 512.5 | 612.5 |
| 150 |  |  |  |  | 612.5 |
| 195 |  |  |  |  | 612.5 |
| 400 |  |  | 432.5 (V) | 512.5 (V) |  |
| 450 |  |  |  |  |  |

- Measure covariance of electron spectrum and photon spectrum

- Request 800 eV, 950 eV and see what the pulse energy (and count rate).

#### Shift 5

With Beam:

- Bring the beam down to the TMO beamline. Note down and point the beam in the same locations of IM2K4 and IM5K4. Day-to-day variations should be noted. (30 mins)

  - Check for KB1 transmission

  - Check for Sample paddle repeatability.

  - See if signal on the SMA tip. (Could try coarse timing at IP)

- Setup energy scan PV. (10 mins). Ask to have a wider range of +/-250 eV.

- Setup FZP at 400 eV. (30 mins)

  - If higher pulse energy > 40 uJ, use Target 2 of the attenuator.

- Spectral Reconstruction scans

  - Increase the transmission to 100% (~200 uJ)

  - Argon L-edge (400 eV)

  - Retardations: 0, 100, 125 V

  - Different hit rate (as a factor of 1, 2, 4, 8, 16* per feature) as a function of needle position

  - Scanning photon energy across the FZP range.

  - 30s at each retardation

- Change sample to NNO.

  - Commission 615 eV FZP.

  - Complete the Shift 4 measurement with FZP.

  - Scan with continuous FZP positioning.

  - Keep the FZP position at the center of the camera for each energy points, and calibrate the YAG-X and Lens-X position.

  - Do a photon energy scan around 400 eV, (390 - 410 eV).

  - Do a pre-edge scan at the Oxygen edge (515 - 535 eV).

| FZP | Nitrogen (410 FZP) |  |  | Si (613) |
| --- | --- | --- | --- | --- |
| Ret/hf | -10 | -5 | 0 | +200 |
| 0 | 402.5 | 405.5 | 412.5 | 612.5 |
| 5 |  |  |  |  |
| 15 |  |  |  |  |
| 45 |  |  |  |  |
| 50 |  |  |  |  |
| 95 |  |  |  |  |
| 100 |  |  |  | 612.5 |
| 150 |  |  |  | 612.5 |
| 195 |  |  |  | 612.5 |

- Find fine timing at the ATM using Target 1b (GaAs).

  - Take 3 minutes of ATM data

  - Repeat for Target 2b

  - Repeat for Target 3b

- Spatial and temporal overlap with 1.3 um at the IP.

  - Align the sample paddle YAG and SiN with the X-ray/laser. Do spatial overlap. Use microscope for this.

  - Measure t0 using the solid target - clear YAG and SiN.

- Measure t0 using NNO.

  - Use photon energy of ~402 eV.

  - Pump-probe with 1.3 um

    - Linear polarization

    - Circular polarization

#### Things we still need to do:

Things that need to be added to a shift plan.

- Sample paddle repeatability.

- More retardation scans.

- NNO measurements (photon energy and retardation)

- Correlation between FZP and ToF.

- Bias - is there more info needed here?

- MRCO Z-scan.

- FZP commissioning - Try 613 eV FZP.

- Establish Downstream focus condition - maybe a low priority

- Check-out ATM (parasitic) - (coarse timing)

  - Look for signals parasitically.

#### Shift 5:

- laser/x-ray spatial overlap on sample paddle

- Laser timing signal on sample paddle

- ATM Setup

- Fine timing in gas phase target (N2)

- Verify X,Y position of XFEL focus: MRCO XY-scan to find the center of beam with respect to the centering ring. 0 retardation. (N2)

- Demonstrate all spectrometers and detectors working with variable retardation

  - set up the retardation and the lens-can of the spectrometers for different retardation

  - Retardation independent spectra for few features at different energy points of N2/NNO.

  - Beta parameters for 0 to full retardation for N2, CO2, Ne.

- Test Current Mode: Ensuring that the spectrometer is tuned for high signal rate mode, and validate MCP performance in cases of increased depletion

  - we need to make sure that the MCP detectors are working at the non-linear gain

  - 5, 10, 25, 50, 75, 100, 150, 200 mV signal level (without amplifier).

  - Repeat 5, 10, 15, 25 after adding amplifiers.

  - (laser only) find the common bias voltage for the MCP. Find the responsivity curve for each MCPs in the wagon-wheel. Measure the current through the MCP. Too many detectors, therefore needs more diagnostics.

- Verify position of XFEL focus: MRCO Z-scan to find nonlinear signal (DCH in N2/NNO/?)

  - If the non-linear is difficult to find, then consider using the blurring effect of space charge to identify the tightest focus.

- Ne/Ar Calibration measurements

  - FZP/MRCO correlation measurements for spectral reconstruction

  - Redo the above with a different hit rate (as a factor of 1, 2, 4, 8, 16* per feature).

  - FZP/MRCO correlation measurements for spectral reconstruction in current mode

- Laser Spatial overlap on paddle

- Laser temporal overlap on paddle (some scope can be shared with X-10096)

  - Verify optimal spatial overlap in same place for paddle vs. gas (SFI of N2)